## Slide 1
CE-OT Management on AI-Software as a Medical Device in Hospitals
Kang-Ping Lin
Chung-Yuan Christian University, Taiwan

## Slide 2: The Team / Workgroup
The Team members include Mei-Feng Chen, Ming-Chin Yang, Chien-Wen Lai, Kang-Ping Lin
The team members of Mei-Feng Chen, and Kang-Ping Lin in Dept. of Electrical Engineering, Chung-Yuan Christian University, work together with Clinical Engineers of Ming-Chih Yang and Chien-Wen Lai in Division of Clinical Engineering, Chang-Hua Christian Hospital, Taiwan for operating the CE-OT Medical Device management.

## Slide 3: Description
Description as noted in the Abstract submission
This presentation intended to propose management practical procedures for AI/ML medical software based on the framework of the full-lifecycle management of the Clinical Engineer (CE) with the Operational Technology (OT) to strengthen with each part of the medical device’s 1) evaluation, 2) procurement, 3) acceptance, 4) management and 5) maintenance, in hospitals.
Chart: Numbers of AI/ML-Enabled Medical Devices
| 2017 | 2018 | 2019 | 2020 | 2021
 | 26 | 61 | 75 | 100 | 38
How to deal with AI/ML-Enabled Medical Devices
via 510(k) clearance, granted De Novo
request, or approved PMA)

## Slide 4: Goals of the project and final users that will benefit
A Clinical Engineer (CE) with operational technology (OT) based on the framework of the full-lifecycle management is responsible for medical device management in hospital. This presentation proposes some CE-OT’s procedure modifications involving into the full-lifecycle management for AI/ML medical software (Software as a Medical Device, SaMD).
CEs may refer these proposed newly added planning items and implement them in actual conditions to achieve the goal of maintaining patient safety and improving the quality of healthcare service.

## Slide 5
EFAI-Bone Age : SaMD Specification Requirement (example)
Software requirements :
Inference Server
− Operating System：Ubuntu 16.04
− Nvidia GPU Drivers：418.56
− Docker：18.09.6
− Tensor RT Inference Server：19.06
− TPM2-Tools：tpm2-tss-2.1.0,
tpm2-abrmd-2.0.3
− PostgreSQL: 12.1
− PACS Software support DICOM
Clinic Computer
− Microsoft Internet Explorer 10 ;
Google Chrome
− TCP/IP Protocols
Hardware Requirements：
Inference Server
− GPU or Display interface：NVIDIA GeForce
GTX1060 6G
− CPU：Intel Core i5-7400 (3.0GHz)
− Memory：DDR4-2400 16GB
− Solid-state disk：128GB
− Hard disk：SATA3 1TB 7200rpm
− Uninterruptible Power Supply (UPS) system
− USB Port：Locked
Clinic Computer
− CPU：Intel Core i5-7400 (3.0GHz)
− Memory：16GB DDR4-2400
− Solid-state disk：128G SSD
− Hard disk：1TB SATA3 7200rpm
− GPU or Display interface：NVIDIA GeForce
GTX1060 6G
− mouse, keyboard
− LCD monitor 17-inch 1366*768 (720p)

## Slide 6
Operational Cooperation
Operational Technology
(Clinical Engineering)
Medical Device
Internet Medical Devices
AI / ML
Medical Devices
Ward
Information Technology
(Computer Center)
PACS
HIS
EHR/EMR Database
Physicians
Nursing station
Nursing station
Clinical Engineer

## Slide 7: Duties of CE Managing on Med. Devices in Hospital
Operational Technology Administrative Structure
( Based on the annual objective in hospital )
Devices Budget
Budget, Non –Budgeting Reviewing Regulation
System: E-budget system
Record: Medical Devices Budget Reviewing Case
Medical Devices Budget Panel Record Improvement: BME department proposal
Procurement Evaluation
Procurement Evaluation Regulation
System: Medical Devices Administrative Information System
Record: specification, assessment table
Improvement: supplier and facility evaluation information
Device Acceptance
Acceptance Regulation
Maintenance Bond Administrative Regulation
System: Medical Device Administrative Information System
Record: attendance record, information, manual, operation introduction, maintenance, form, cycle, Electrical safety, test record
Device Maintenance and Repairing
Maintenance & Repairing Application Regulations
System: Medical Device Administrative Information System
Record: Repairing and Maintenance Record, Maintenance Schedule
Improvement: KPI, Case
Device Unused and Discard
Unused and Discard Regulations
Record: Discard Record
SaMD
SaMD
SaMD

## Slide 8
Medical equipment asset inventory and risk assessment
Medical device Identification
Asset management inventory
Register an inventory of critical infrastructure assets
Identify the risks of availability & impact
Report the risks
Hospital management and crisis management measures
Internet connection
No internet connection, or
can be connected but not connected yet
SaMD

## Slide 9: Results
Management Structure
Strategic
Management
Technical
Applicability :
The scope of the cybersecurity management system
Evaluation :
Status of resource invest in cybersecurity
Cybersecurity:
Operational planning and implementation
Asset management
Risk management
Outsourcing management
Electronic data protection
Communication and operation safety
Adverse event notification and handling
Cybersecurity system development and maintenance

## Slide 10: Results
Checklist for technical Inspection (1~7)
 | Inspection Item | Test | % Score
1 | User’s computer (IP) cybersecurity Inspection | Vulnerability scan test | 10
 |  | Cybersecurity protection system test | 20
2 | Internet malicious activity Inspection | Ability test for blocking malicious relay station attacks | 5
3 | Core cybersecurity system Inspection | Penetration test for internal network | 20
 |  | Protection benchmark test | 5
4 | Internet structure system Inspection | Structure test | 10
5 | Active directory (AD) cybersecurity Inspection | AD cybersecurity test | 5

## Slide 11: Results
Checklist for technical Inspection (1~7)
 | Inspection Item | Test | % Score
6 | IoT connected equipment cybersecurity Inspection | Webcam test | 10
 |  | Access control equipment test | 
 |  | Network printer test | 
 |  | Wireless route/station test | 
 |  | Environmental control system test | 
7 | Configuration setting cybersecurity Inspection | Configuration test for operation system | 15
 |  | Configuration test for Browser | 
 |  | Configuration test for Network communication equipment | 
 |  | Configuration test for APP | 
Total |  |  | 100

## Slide 12: Conclusion
Software as a medical device (SaMD) will gradually involve various sections for clinical use, including physicians, nurses, medical technicians, etc.
Clinical engineers will play an important role on the AI/ML based medical device operations and information integration.
A practical and complete planning and management will help operations of the internet connected software in medical system and improve patient safety and healthcare quality

## Slide 13: Kang-Ping Lin
kplin@cycu.edu.tw
Chung-Yuan Christian University, Taiwan